## Slide 1: Effect on Library Size on Protein Identification and Quantification: TimsTOF
Advanced Clinical Biosystems Research Institute
Qin Fu
Niveda Sundararaman

## Slide 2
Trapped Ion Mobility Spectrometry Time-of-Flight Mass Spectrometry
(Tims-TOF)
Fast, reproducible and multi-dimensional
Retention time
Ion mobility
Mass
Signal intensity
4-dimensional (4D) features tims TOF pro2
Time
Intensity
Time
M/Z
Mobility (1/Ko)
m/z
Sample A. 100% HeLa
Sample B. 50% HeLa and 50% SILAC HeLa
Sample C. 10% HeLa and 90% SILAC HeLa
Sample D. 5% HeLa and 95% SILAC HeLa
Sample E. 100 % SILAC Hela
Samples description
Notes: trapped ion mobility spectrometry-time-of-flight mass spectrometry (TIMS-TOF)

## Slide 3
Real-time Search Results with PaSER
Acquisition
Real time
database search
Peptide number, proteins number
(QCs, blanks and samples)
Continue
Acquisition
or Stop
Real time decision
Data monitoring:
1. Real time search peptides number and proteins number
2. Intensity vs time (Retention time stability)
3. Mobility vs m/z
4. m/z vs time
5. mass accuracy
6. FWHM.Scans (full width at half maximum Scans, data point across the peak)

## Slide 4
LC MS/MS Performance Monitoring
Intensity vs time (Chromatograms and retention time consistence-25 injections of Hela A to E)
Mobility vs m/z
m/z vs Time

## Slide 5
Loading Optimization
Total precursors vs reproducibility vs no missingness
No missingness and reproducibility
Loading for the study
400 ng

## Slide 6
Library Generation
Method 1-10XDDA runs with Hela sample (A, 100% Hela), small library
Method 2-Combine 10xDDA runs library with Bruker’s human library, big library
 | Small Library | Big Library
Proteins | 3926 | 13117
Peptides | 26670 | 350788
Modified Peptides | 32246 | 516016
PrecursorIds | 37373 | 576624

## Slide 7
Search Results
Small Library- A, B, C, D, E group (5 injections) searched with Paser/Tims Diann
Big Library- A, B, C, D, E group (5 injections) searched with Paser/Tims Diann
 | Small Library Results | Big Library Results
Proteins | 3223 | 5245
Peptides | 18689 | 28938
Modified Peptides | 21323 | 39701
Precursors | 23888 | 43091
Proteins

## Slide 8
Filter criteria:
Present in A, B, C, D groups
Present in all injections (no missingness, 20 out of 20 injections)
CV<20% within the group (from 5 injections)
Sample groups: A (100%), B (50%), C (10%) and D (5%)
 | Small Library Results | Big Library Results
Proteins | 486 | 530
Peptides | 1096 | 1130
Modified Peptides | 1146 | 1301
Precursors | 1182 | 1334
Proteins
Narrow Search Results with Filters

## Slide 9
Lower limit of Quantification
Chart: 
| 100 | 50 | 25 | 12.5 | 6.25 | 3.125 | 1.5625
 | 100 | 50 | 25 | 12.5 | 6.25 | 3.125 | 1.5625
Undiluted Analytes pool
2-5-fold dilutions with analyte-free matrix
N= 5 of each level
Performed twice
Until analyte no longer quantifiable
CV and Bias (~recovery) compared to expected values
Grant and Hoofnagle, Clin Chem 2014

## Slide 10: Demo: Count Number of Precursors per Concentration Category
The larger library shows higher number of precursors across each category
Small Library
Large Library

## Slide 11: Demo: CV Calculation Across Injections within each Concentration Category
In both cases, higher concentrations of Hela (left: A) shows the lowest CVs, with gradual increase in CV as concentration decreases.
Small Library
Large Library

## Slide 12: Demo: Precursor Abundance across Concentration Categories
Small Library
Large Library
Even though the big library shows higher number of precursors and proteins, the smaller library shows tighter trends towards the expected patterns.

## Slide 13: Demo: RT Difference between Precursors
Small Library
Large Library
RT difference between common precursors across the concentration categories indicates that the difference window is slightly narrower in the case of the small library

## Slide 14: Demo: Ion Mobility(IM) Difference between Precursors
Small Library
Large Library
IM difference between common precursors across the concentration categories

## Slide 15: Demo: Distribution of Precursors based on Correlation b/w Measured and Expected Trends
Good:
Precursors with CV < 20% and no missingness
Correlation b/w measured and expected trend >= 0.9
Bad:
Precursors with CV < 20% and no missingness
Correlation b/w measured and expected trend b/w 0.5 – 0.7
Ugly:
Precursors with CV < 20% and no missingness
Correlation b/w measured and expected trend <= 0.4

## Slide 16
Precursors with Good Regression

## Slide 17
Precursors with Bad Regression

## Slide 18: Conclusions
Large library has more precursors and proteins compared to small library
Search results also yield the same trend when comparing large library to small library
CV trends indicate that the higher concentrations of Hela have smaller variations across injections whereas lower concentrations show higher variations.
Larger does not always mean better i.e., search via the large library yields higher number of precursors but the linearity trend of the search using the smaller library is closer to expected trend.

## Slide 19: Access to Input Files/Scripts/Results
Instructions for Downloads/Setup Required before ASMS Workshop located on GitHub Link:
https://github.com/csmc-vaneykjlab/ASMS-Workshop-TimsTOF-Analysis
Link to Input Files, Results, Script for visualization and Setup Guidelines provided.